- General

  - Manufacture Background

    - All magmeters furnished shall be manufactured by a registered ISO 9001 quality standard facility

    - All meters are to be manufactured in the USA

    - Meters supplied shall be from a US based company that has:

      - Manufactured water meters for at least twenty five (25) years

      - Distributes all types and size magnetic flowmeters as indicated in these specifications

  - Type

    - Meters shall be of the magnetic flow meter type, all of which have NSF compliance

- Capacity

  - Three Inch Meter

    - Normal Operating Range: 3.62 - 723 gpm

    - Maximum Continuous Flow: 723 gpm

    - Maximum Loss of Head at Max Cont. Flow: <1 psi

  - Four Inch Meter

    - Normal Operating Range: 6.43 - 1285 gpm

    - Maximum Continuous Flow: 1285 gpm

    - Maximum Loss of Head at Max Cont. Flow: <1 psi

  - Six Inch Meter

    - Normal Operating Range: 14.46 - 2891 gpm

    - Maximum Continuous Flow: 2891 gpm

    - Maximum Loss of Head at Max Cont. Flow: <1 psi

  - Eight Inch Meter

    - Normal Operating Range: 25.70 - 5140 gpm

    - Maximum Continuous Flow: 5140 gpm

    - Maximum Loss of Head at Max Cont. Flow: <1 psi

  - Ten Inch Meter

    - Normal Operating Range: 40.15 - 8031 gpm

    - Maximum Continuous Flow: 8031 gpm

    - Maximum Loss of Head at Max Cont. Flow: <1 psi

  - Twelve Inch Meter

    - Normal Operating Range: 57.82 - 11565 gpm

    - Maximum Continuous Flow: 11565 gpm

    - Maximum Loss of Head at Max Cont. Flow: <1 psi

- Size

  - Overview

    - Determined by the nominal size (in inches) of the opening in the inlet and outlet flanges

  - Specifications

    - Three Inch Meter Lay Length: 12.25”

    - Four Inch Meter Lay Length: 10.24”

    - Six Inch Meter Lay Length: 12.27"

    - Eight Inch Meter Lay Length: 14.24”

    - Ten Inch Meter Lay Length: 18.18”

    - Twelve Inch Meter Lay Length: 19.68”

- Product Details

  - Functions/Performance

    - Operating Temperature: 10° to 130° F (-12° to 54° C)

    - Storage Temperature: -40° to 158° F (-40° to 70° C)

    - EMI/RFI protection: Per EN 61326-1:2013

    - Pressure rating: Suitable for use in water systems with up to 150 psi normal operating pressure

    - Diagnostics: Self diagnostics with on screen display of faults.

    - Rate Display: Field configurable digital indicator displaying flow in liters, cubic meters, cubic feet, and gallons.

    - Flow Totalizer

      - A fully configurable totalizer integral to the transmitter

      - Totalized flow shall be displayed

      - Flow shall be capable of being displayed in forward total, reverse total, and net total

    - Empty Pipe Performance: Meter shall be capable of running empty indefinitely without damage to any component

    - Empty Tube Zero

      - The transmitter shall lock the output at zero when no flow is detected

      - The empty tube zero feature shall be enabled automatically when the transmitter detects no flow

    - Low Flow Cutoff: The transmitter shall automatically drop the flow rate displayed and outputs to zero when the flow rate is below the minimum rated flow range for that meter size.

    - Data Logger

      - The meter shall have the ability to add a data logger to collect data up to two years

      - Data intervals as low as 15 seconds

      - Data storage option to stop the collection of data when memory is full or automatically wrap around

    - Low Battery Warning: The meter shall display a low battery warning when the batteries are running low

    - All meters shall meet or exceed AWWA C701 performance standards

  - Physical

    - Metering Spool:

      - Epoxy-coated welded steel (4”-12”)

    - Powder coated duxctile iron (3”)

    - Flanges: 150 lb. ANSI pattern

    - Liner

      - Santoprene flange/Polypropylene liner body (4”-12”)

      - Noryl® (3”)

    - Electrodes: 316 stainless steel standard

    - Housing: Meters shall be constructed of powder coated die-cast aluminum and designed to meet NEMA 6P (IP68) environmental conditions

    - Finish: All external surfaces shall have a chemical and corrosion resistant finish

    - Grounding: Equalization lugs will be provided as standard

    - Transmitter shall be integral to meter

  - Power

    - Must have battery backup for power outage

  - Accessories/Documentation

    - Factory Calibration: All meters shall be factory calibrated and a copy of the report for each meter shall be furnished with the meter

  - Connections

    - Meter body shall be flanged

    - 3” through 12” sizes shall be round flanged 150# ANSI pattern

- Flow Converter Details

  - Type

    - Microprocessor-based, intelligent transmitter shall be mounted integral to, and on top of, the flow tube

    - The flow converter / transmitter housing shall be powder-coated die-cast aluminum

  - Display

    - Display shall be sealed, indicating flow rate and totaling any combination of cubic feet, gallons, liters, or cubic meters

    - Display shall be integral to the converter/transmitter

  - Transmitter Box Sealing

    - The flow converter/transmitter cover shall have a tamper-proof seal that must be destroyed in order to remove the cover

  - Meter Serial Number

    - The meter serial number shall be imprinted on the flow converter

- Registration Accuracy

  - Registration accuracy

    - Over the normal operating range shall be 99% to 101%

    - At low flow shall not be less than 98%

- Output Capability

  - All meters:

    - Equipped with an isolated current sinking and user scalable pulse output

    - Meet or exceed AWWA C701 performance standards

Acceptable meters shall be Seametrics iMAG 4700 or approved equal.

LT-14358r1.0 20170727

7/27/17